Abstract:Objectives: Different techniques and choices exist for revision of pedicle screws, two of which are pedicle screw combined with cement augmentation and expansive pedicle screw &amp;#64257;xation. This biomechanical study was designed to compare the pullout strengths of two different revision techniques. Methods: Fourteen lumbar vertebrae obtained from four calves (mean age 15 months) were divided into two groups equal in number. Monoaxial 6.0-mm pedicle screws were inserted into the right pedicles, and axial pullout testing was performed at a rate of 10 mm/min and failure strengths were recorded. Revision was performed with the same-sized pedicle screws reinforced with polymethylmethacrylate in group 1, and with 7.0-mm expansive pedicle screws in group 2, and pullout testing was repeated to record maximum revision pullout strengths. Results: The mean pullout strengths recorded before and after revision were signi&amp;#64257;cantly different in both groups, being 2,162.9&plusmn;718.5 N and 2,794.3&plusmn;979.2 N in group 1 (p=0.041) and 2,605.0&plusmn;487.6 N and 3,327.1&plusmn;640.8 N in group 2 (p=0.012), respectively. However, the mean pullout strengths recorded before and after revision did not differ signi&amp;#64257;cantly between the two groups (p&gt;0.05). Conclusion: Our results showed that expansive pedicle screws 1 mm larger in diameter provide similar pullout strengths to those of same-sized, polymethylmethacrylatereinforced screws in revision of pedicle screw &amp;#64257;xation, suggesting that they can be preferred with the additional advantages of ease of application and avoiding risks for pedicle fracture and cement leakage.